## Sheet3 (hidden)
Photon energy | Electron energy
4
6
8
10
15
18
6 FFF
7FFF
8FFF
10FFF
11FFF

## Sheet1
DO NOT MODIFY THE WORK SHEET. FILL UP THE BLUE CELLS ONLY
Name of the Medical Physicist responsible for the measurements
Date of completion of the measurements
1 | General Information of the Linear Accelerator unit
1.1 | Manufacturer
1.2 | Model
1.3 | Sr. No.
1.4 | Normal treatment distance (NTD)
1.5 | Field size (Minimum in cm 2) | [7 empty] | Photon(s) only
1.6 | Field size (Maximum in cm 2) | [7 empty] | Photon(s) & Electrons
1.7 | Electron applicator(s) field sizes, if applicable
1.8 | Radiation beam type ( Photon Only/Photon(s) and Electron (s))
1.8.1 | Nominal flattened photon beam energies (MV) | [7 empty] | Available
1..8.2 | Nominal un-flat photon beam energies(MV) | [7 empty] | Not available
1.8.3 | Nominal electron beam energies (MeV)
1.8.4 | Dose rate ranges (flattened photon beam energies)
1.8.5 | Dose rate ranges (un-flat photon beam energies) | [7 empty] | Secondary
1.8.6 | Dose rate ranges (Electron beam energies) | [7 empty] | Tertiary
1.9 | Available Treatment options (Conventional /3D CRT/ IMRT/ IGRT/SRS/SRT/SBRT/VMAT/ Any Other)
1.1 | Available imaging modality and energy (e.g. Portal Imaging, kVCT, MVCT etc.) | [7 empty] | Yes
1.10.1 | Specify salient technical features of Imager (such as: EPID, Flat Panel, FOV, any other) | [7 empty] | No
1.11 | Any other ionizing radiation based imaging/positioning system
1.12 | Multileaf Collimator (MLC) ( Available/Not available)
1.12.1 | Type of in-built MLC ( Secondary/ Tertiary)
1.12.1.1 | Make and Model
1.12.1.2 | Minimum definable field size at isocentre
1.12.1.3 | Maximum definable field size at isocentre
1.12.1.4 | Number of leaves
1.12.1.5 | Projected leaf width at isocenter
1.12.1.6 | Over travel distance
1.12.2 | Details of add-on MLC
1.12.2.1 | Make and Model
1.12.2.2 | Minimum definable field size at isocentre
1.12.2.3 | Maximum definable field size at isocentre
1.12.2.4 | Number of leaves
1.12.2.5 | Projected leaf width at isocenter
1.12.2.6 | Over travel distance
1.13 | Couch type (e.g. 4D/6D) and material
1.14 | Any other technical information
2 | Performance test results of the Conventional/3DCRT Capable LINAC
Sr. No. | Parameters | Tolerance Limit | Observation ( Write “Yes/No or give values of the measurement”, as the case may be) | Observation is meeting the requirments | Remarks, if any
2.1 | Mechanical Tests ( Couch)
2.1.1 | Minimum Level of table top above floor | ≤ 80 cm
2.1.2 | Longitudinal motion of couch | ≥ 70 cm
2.1.3 | Vertical motion of couch from isocentre | ˃40 cm below isocentre and ˃3 cm above isocentre
2.1.4 | Lateral motion of couch | ± 20 cm
2.1.5 | Rotational motion of couch (----deg to -----deg) | At least ±90 degree
2.1.6 | Minimum Linear speed of table top | Min ≤ 1 cm/sec
2.1.7 | Maximum Linear speed of table top | Max ≤ 10 cm/sec
2.1.8 | Minimum Rotational speed of couch | Min: ≤ 1°/sec
2.1.9 | Maximum Rotational speed of couch | Max: ≤ 7°/sec
2.1.10 | Accuracy of longitudinal motion of couch/ table top | 0.3 cm for speed ≤2.5 cm /sec & 1 cm for > 2.5 cm /sec
2.1.11 | Accuracy of Lateral motion of couch/ table top | 0.3 cm for speed ≤2.5 cm /sec & 1 cm for > 2.5 cm /sec
2.1.12 | Accuracy of Vertical motion of couch/ table top | 0.3 cm for speed ≤2.5 cm /sec & 1 cm for > 2.5 cm /sec
2.1.13 | Accuracy of angular motion of couch | 0.5° at ≤ 1°/sec & 3° for > 1°/sec
2.1.14 | Table top sag at isocenter when loaded with 135kg distributed over 2m through isocenter | ≤ 2 mm
2.1.15 | Shift in centre of optical field due to vertical motion of couch from min. to max. position | ≤ 2 mm
2.1.16 | Shift in Isocentre due to couch rotation | ≤ 2 mm dia
2.1.17 | Couch Transmission | Provide Manufacturer’s specified Tolerance______
2.2 | Mechanical Tests (Collimator)
2.2.1 | Rotational Motion of collimator (----deg to -----deg) | ≥ ±100°
2.2.2 | Accuracy of angular scale of collimator | 0.5 deg at ≤ 1°/sec and 3 deg for > 1°/sec
2.2.3 | Gap between isocentre and end of accessory mount | ≥ 30 cm
2.2.4 | Shift in Isocentre due to Collimator rotation | ≤ 2 mm dia.
2.2.5 | Minimum Rotational speed of collimator | Min ≤ 1°/sec
2.2.6 | Maximum Rotational speed of collimator | Max ≤ 7°/sec
2.2.7 | Minimum Jaw Speed | Min: ≤ 1 cm/sec
2.2.8 | Maximum Jaw Speed | Max: ≤ 10 cm/sec
2.2.9 | Orthogonality of the adjacent jaws | ±1°
2.2.10 | Symmetry of the jaws | ±1 mm
2.2.11 | Optical field intensity | ≥ 40 Lux
2.2.12 | Coincidence of MLC and secondary collimator axes | 1 mm (applicable for tertiary and add on MLC only)
2.2.13 | Reproducibility of leaf positions | ±1 mm
2.2.14 | Accuracy of Leaf positions | ±1 mm
2.2.15 | Congruence between MLC fields and beam’s eye view (BEV) | ±2 mm (This test should be carried out for regular and irregular fields)
2.2.16 | Accuracy of optical field size in the range of 5x5 to 40x40 cm2 (symmetrical field) | (i) ±1 mm for Fs ≤ 10x10 cm2 (ii) ±1% (not exceeding 2mm) of Fs, for Fs> 10x10 cm2
2.2.17 | Accuracy of optical field size for a range of asymmetric field | (i) ±1 mm for Fs ≤ 10x10 cm2 (ii) ±1% (not exceeding 2mm) of Fs, for Fs> 10x10 cm2
2.3 | Mechanical Tests (Gantry)
2.3.1 | Accuracy of gantry angle calibration | ≤ 0.5°
2.3.2 | Accuracy of gantry angular position during arc motion | 0.5° for ≤ 1°/sec and 3° for > 1°/sec
2.3.3 | Shift in isocentre due to gantry rotation | 2 mm dia.
2.4 | Other mechanical tests
2.4.1 | Coincidence among the Mechanical front pointer, optical distance indicator (ODI) and lasers | 2 mm dia. sphere
2.4.2 | Accuracy of ODI in the range of ±20 cm from isocentre | ± 2 mm
2.4.3 | Accuracy of laser alignment with isocentre | ± 2 mm
2.4.4 | Positional reproducibility of physical wedges | 2 mm or 2% change in wedge transmission
2.4.5 | Positional reproducibility for Tray | 2 mm
2.4.6 | Positional reproducibility for electron applicator | 2 mm
2.4.7 | Latching of wedges and blocking tray provided and working
2.5 | Electrical test
2.5.1 | Treatment room door interlock provided and working
2.5.2 | Emergency OFF switches provided and working
2.5.3 | Interrupt switch provided and working
2.5.4 | Anti-collision system provided and working
2.5.5 | Whether all movements of the table and accelerator are controlled using the pendant
2.5.6 | Are all operating parameters displayed on the console/in room monitor and functional?
2.5.7 | Are safety switches operable from control console?
2.5.8 | Is arc / rotation switch start / terminate the rotation at the pre-assigned angle?
2.5.9 | Is mode selection switch/option provided?
2.5.10 | Is irradiation feasible only after mode selection?
2.5.11 | Does alteration of the mode during the irradiation cuts off the beam
2.5.12 | Does primary dose monitoring system of a primary/secondary dose monitoring system terminates irradiation when the selected number of dose monitor units has been reached
2.5.13 | Is secondary dose monitoring system terminates irradiation when the pre-selected number of dose monitor units has been exceeded, either by not more than l0% if a percentage margin is used, or by not more than the equivalent of 0.25 Gy absorbed dose at NTD
2.5.14 | Is irradiation possible only after having pre-selected monitor units on the control console?
2.5.15 | Is the number of monitor units continued to be displayed until reset to zero?
2.5.16 | Are residual monitor units displayed (at least for 20 min.) even after accidental or intentional power cut off?
2.5.17 | Are dose monitor systems functions independently?
2.5.18 | Is each dose monitor system capable of terminating the irradiation independently?
2.5.19 | Is independent timer provided?
2.5.20 | Does the timer start and stop simultaneously with the irradiation?
2.5.21 | Is wedge filter integrated into unit head with pre-selected mode of insertion and interlocked with console?
2.5.22 | Whether angle/ orientation of the wedge filter displayed during use?
2.5.23 | Is shielding tray interlock provided and functional?
2.5.24 | Is it true that irradiation in photon mode is not possible, if accessories of electron mode are in use and vice-versa?
2.5.25 | Is radiation type displayed on control panel before, during and after irradiation?
2.5.26 | Is range of photon and electron energies displayed on panel?
2.5.27 | Is selected energy displayed on panel during and after irradiation?
2.5.28 | Whether no irradiation possible without wedge filter selection on control panel?
2.6 | Dosimetry Test (Photon Beam)
2.6.1 | Congruence between optical and radiation fields (Symmetric Fields:5x5 cm2, 10x10 cm2, 20x20 cm2, 30x30 cm2) | ± 2 mm
2.6.2 | Congruence between optical and radiation fields (Asymmetric fields defined by over travel option of the collimator) for most useful clinically photon beam and for field sizes mentioned below: 10(X1= - 5cm and X2= 15cm)*10 (Y)cm 10 (X) cm*10(Y1= -5 cm and Y2= 15 cm | ± 2 mm
2.6.3 | Overlapping between parallel opposed radiation fields defined by jaws for 10x10 cm2 for most useful clinically photon beam | 3 mm
2.6.4 | Overlapping between parallel opposed radiation fields defined by MLC for 10x10 cm2 for most useful clinically photon beam | 3 mm
2.6.5 | Accuracy of radiation isocentre | 2 mm dia sphere
2.6.6 | Flatness for symmetrical fields at 10 cm depth in isocentric set-up for all available photon beam energies and for field sizes 5x5, 10x10 and 30x30 cm2 | Flatness for field sizes 5x5 cm2 to 30x30 cm2 : ≤ 106 % & for field size >30x30 cm2: < 110 %
2.6.7 | Symmetry for symmetrical fields at 10 cm depth in isocentric set-up for all available photon beam energies and for field sizes 5x5, 10x10 and 30x30 cm2 | ≤ 103 %
2.6.8 | Radiation beam penumbra for 10x10 cm2 at dmax for all available photon beam energies | ≤10 mm
2.6.9 | Output (cGy/MU) for all available photon beam energies for field size 10x10 cm2 under reference condition
2.6.10 | Output Constancy at different times in a day for all available photon beam energies and for field size 10x10 cm2 | 2%
2.6.11 | Output constancy at cardinal positions (00, 900, 2700 & 1800) of the gantry for all available photon beam energies and for field size 10x10 cm2 | 2%
2.6.12 | dmax for available photon energies | Typical value with possible variations: 1.0±0.2 cm for 4 MV 1.5±0.2 cm for 6 MV 2.3±0.2 cm for 10MV 2.9±0.2 cm for 15 MV 3.2± 0.2 cm for 18 MV
2.6.13 | PDD for 10x10 cm2 at 10 cm depth for available photon energies | Typical value with possible variations: 63.0%±2% for 4 MV 67.5%±2% for 6 MV 73.0%±2% for 10MV 77.0%±2.0% for 15 MV 79.0%±2.0% for18MV
2.6.14 | Percent surface dose for 30x30 cm2 [(Ds=D0.5mm/Ddmax)*100] for available photon energies 4 to 18 MV) | ≤ 70% for 4 MV ≤ 60% for 6 MV ≤ 50% for 8 to 18 MV
2.6.15 | Quality Index (TPR20/10) for available photon energies | Typical value: 0.626±3% for 4 MV 0.677±3% for 6 MV 0.731±3% for 10MV 0.763±3% for 15 MV 0.775±3% for 18MV
2.6.16 | Energy stability (quality index) for available photon energies at different times in a day | ± 1%
2.7 | Dosimetry Test (Electron Beam)
2.7.1 | Congruence between optical and radiation fields for 10x10 cm2 applicator and for most useful clinically electron beam | ± 2 mm
2.7.2 | Flatness along major axes at the standard measurement depth (as specified in IEC standard) as separation between 90% and 50% dose points on either side of the beam profile for all available electron beam energies: i)10x10 cm2 applicator ii) Maximum applicator size | ≤ 10 mm
2.7.3 | Symmetry along major axes at depth of dose maxima in SSD set-up (maximum ratio of absorbed doses at symmetrical points from central beam axis and more than 1 cm inside the 90% isodose contour ) for all available electron beam energies | 105%
2.7.4 | Radiation field penumbra at dmax for 10x10 cm2 applicator size for all available energies | ≤15 mm
2.7.5 | dmax for 10x10 cm2 applicator size for following energies in MeV | ±2 mm of the manufacturer specified value |  |  | Manufacturer specified value…….
2.7.6 | X-ray contamination for all available energies for 10x10 cm2 applicator size | <4% up to 6 MeV, <5% for 6-15 MeV, <10% for > 15 MeV
2.7.7 | Most probable energy Ep,0 at phantom surface for 20x20 cm2 applicator size for all available energies | Quote the calculated values from beam data
2.7.8 | R50 (cm) in water for all available energies for 10x10 cm2 applicator size | Provide Manufacturer’s specified value |  |  | Manufacturer specified value…….
2.7.9 | Energy stability for 10x10 cm2 applicator in terms of R50 for all available energies at different times in a day | ± 4%
2.7.10 | Output calibration (cGy/MU) for all available electron beam energies using applicator of 10x10 cm2 in SSD set-up
2.7.11 | Output Constancy at different times in a day for all available electron beam energies and for 10x10 cm2 applicator size | ± 2%
2.8 | Dose Monitoring System
2.8.1 | Reproducibility for most useful clinically photon beam and most useful clinically electron beam for 10x10 cm2 applicator size | CoV ≤ 0.5% | CoV for X: CoV for E:
2.8.2 | Linearity for most useful clinically photon beam and most useful clinically electron beam for 10x10 cm2 applicator size (for ≥ 5 MU ) | CoL≤ 2 % | Linearity for X: Linearity for E:
2.8.3 | Dependence of monitor response on gantry rotation for most useful clinically photon beam and most useful clinically electron beams for 10x10 cm2 applicator size | ≤ 3%
2.8.4 | Temporal stability of the monitor response at different times of the day for most useful clinically photon beam and most useful clinically electron beams for 10x10 cm2 applicator size | ≤ 2%
2.9 | Radiation Safety
2.9.1 | *Maximum Photon leakage radiation through beam limiting devices (Secondary collimator) at NTD | Maximum:2%
2.9.2 | *Average Photon leakage radiation through beam limiting devices (Secondary collimator) at NTD | Average : 0.75%
2.9.3 | *Maximum Photon leakage radiation through MLC when either pair of jaws is replaced with MLC | Maximum: 2%
2.9.4 | *Average Photon leakage radiation through MLC when either pair of jaws is replaced with MLC | Average : 0.75%
2.9.5 | *Photon leakage radiation through MLC when MLC is used as tertiary jaws | 5% (Maximum)
2.9.6 | *Maximum Photon leakage radiation in the patient plane# | Maximum : 0.2%
2.9.7 | *Average Photon leakage radiation in the patient plane# | Average : 0.1%
2.9.8 | *Maximum Neutron leakage radiation in the patient plane# | Maximum : 0.05%
2.9.9 | *Average Neutron leakage radiation in the patient plane# | Average : 0.02 %
2.9.10 | *Maximum percentage of photon leakage radiation at 1 m from the target; path of electrons between electron gun and the target or electron window in other than patient plane | 0.5% ( Photon)
2.9.11 | *Maximum percentage of neutron leakage radiation at 1 m from the target; path of electrons between electron gun and the target or electron window in other than patient plane | 0.05% (Neutron)
2.9.12 | The ambient dose equivalent due to induced activity in the components of the accelerator$ at 5 cm from the surface | 200 µSv/h
2.9.13 | The ambient dose equivalent due to induced activity in the components of the accelerator$ at 1 m from the surface | 20 µSv/h
*Percentage leakage radiation is percentage of maximum absorbed dose for 10 cm x 10 cm radiation field size measured on central beam axis at NTD # Circular plane of radius 2 m centred on radiation beam axis at NTD $As per AERB safety code No. AERB/RF-MED/SC-1 (Rev. 1)
3 | Performance test results of the Filtered Beam IMRT Linear Accelerator (Note: Following tests need to be carried out for IMRT capable LINAC in addition to the tests for conventional LINAC)
Sr. No. | Parameters | Tolerance Limit | Observation ( Write “Yes/No or give values of the measurement”, as the case may be) | Observation is meeting the requirments | Remarks, if any
3.1 | Mechanical Tests
3.1.1 | Positional reproducibility of longitudinal motion of the couch | ± 2 mm
3.1.2 | Positional reproducibility of lateral motion of the couch | ± 2 mm
3.1.3 | Positional reproducibility of Vertical motion of the couch | ± 2 mm
3.1.4 | Shift in centre of optical field with collimator axis of rotation | 1 mm
3.1.5 | Laser alignment with isocentre | ±1.5 mm
3.1.6 | Leaf position accuracy (at 0° and 90° /270° gantry positions) | ±1 mm
3.1.7 | Leaf position reproducibility (at 0° and 90° /270° gantry positions) | ± 1 mm
3.1.8 | Accuracy of Leaf speed (at 0° and 90° /270° gantry positions) | ± 0.5 cm/s
3.2 | Dosimetry Tests
3.2.1 | Output consistency for low MU settings (for 2-4 MU) at nominal dose rate for 6 MV photon beam and for field size 10x10 cm2 | CoV ≤ 5%
3.2.2 | Output consistency at lowest and highest dose rate settings for 6 MV photon beam and for field size 10x10 cm2 | ≤ 2%
3.2.3 | Output consistency at 0° and 90° /270° gantry positions for 6 MV photon beam and for field size 10x10 cm2 | ± 1%
3.2.4 | Beam flatness consistency for field size 5x5 cm2 at 10cm depth for 6 MV photon beam | ± 1%
3.2.5 | Beam symmetry consistency for field size 5x5 cm2 at 10cm depth for 6 MV photon beam | ± 1%
3.2.6 | Consistency of PDD (10x10 cm2 at 10 cm depth) for 6 MV photon beam | 1%
3.2.7 | Consistency of TMR/ TPR (10x10 cm2 at a typical clinical depth) for 6 MV photon beam | 1%
3.2.8 | Leaf positioning accuracy e.g. Garden fence test at different MLC speeds and dose rate for 6 MV photon beam | 0.5 mm
4 | Performance test results of the SRS/SRT/SBRT LINAC (Note: Following tests need to be carried out for SRS/SRT/SBRT capable LINAC in addition to the tests for conventional LINAC)
Sr. No. | Parameters | Tolerance Limit | Observation ( Write “Yes/No or give values of the measurement”, as the case may be) | Observation is meeting the requirments | Remarks, if any
4.1 | Mechanical Tests
4.1.1 | Accuracy of field size indicator with MLC used for SRS/SRT/SBRT (for field size up to 5x5 cm2) | ± 1 mm
4.1.2 | Accuracy of Laser alignment with isocentre | ± 1mm
4.1.3 | Shift in centre of optical field with collimator axis of rotation | ± 1 mm
4.1.4 | Shift in centre of optical field with couch axis of rotation | ± 1 mm
4.1.5 | Shift in centre of optical field with gantry axis of rotation | ± 1 mm
4.1.6 | Accuracy of longitudinal position indicators of the couch | ± 1 mm
4.1.7 | Accuracy of lateral position indicators of the couch | ± 1 mm
4.1.8 | Accuracy of vertical position indicators of the couch | ± 1 mm
4.1.9 | Accuracy of angular position indicator of the couch | 0.5°
4.2 | Dosimetry Tests
4.2.1 | Radiation isocenter accuracy e.g. Winston-Lutz test | 1mm dia sphere
4.2.2 | Dependence of monitor response on gantry rotation for 6 MV photon beam and for field size 3 x3 cm2 | ± 1%
4.2.3 | Beam flatness constancy for field size 3x3 cm2 at 10cm depth for 6 MV photon beam | ± 1%
4.2.4 | Beam symmetry constancy for field size 3x3 cm2 at 10cm depth for 6 MV photon beam | ± 1%
4.2.5 | Constancy of x-ray beam quality for 6 MV photon beam and for field size 3 x3 cm2 | 1%
4.2.6 | Monitor unit linearity for 6 MV photon beam and for field size 3 x3 cm2 | ± 2% for MU ≥ 5, ± 5% for 2 - 4 MU
4.2.7 | Reproducibility of PDD (10x10 cm2 at 10 cm depth) for 6 MV photon beam | 1%
4.2.8 | Reproducibility of TMR/ TPR (for 3x3 and 10x10 cm2 at a typical clinical depth) for 6 MV photon beam | 1%
4.2.9 | Reproducibility of output factors (SRS/SRT/SBRT for field sizes 1x1, 2x2, 3x3, 4x4 and 5x5 cm2) for 6 MV photon beam | 1%
5 | Performance test results of the Volumetric Modulated Arc Therapy (VMAT)@ Linear Accelerator (Note: Following tests need to be carried out for VMAT capable LINAC in addition to the tests for conventional LINAC + IMRT capable LINAC)
Sr. No. | Parameters | Tolerance Limit | Observation ( Write “Yes/No or give values of the measurement”, as the case may be) | Observation is meeting the requirments | Remarks, if any
5.1 | Mechanical Tests
5.1.1 | Accuracy of gantry angular position during arc motion | 1° for speed >1°/sec
5.1.2 | Maximum Leaf Speed | Manufacturer’s specified tolerance |  |  | Manufacturer specified value…….
5.1.3 | Maximum Gantry Speed | Manufacturer’s specified tolerance |  |  | Manufacturer specified value…….
5.1.4 | Accuracy of Leaf Speed | ≤ 2%
5.1.5 | Accuracy of leaves position | <0.5 mm
5.2 | Dosimetry Tests
5.2.1 | Monitor chamber constancy during Arc mode of the gantry | 2%
5.2.2 | Output constancy with varying dose rate and full Arc motion of the gantry | ≤ 2%
5.2.3 | Accuracy of dose with variation of dose rate and gantry speed | ≤ 2%
5.2.4 | Beam flatness over full Arc motion of the gantry; detector attached to gantry head | ≤ 2% difference with respect to static beam flatness
5.2.5 | Beam flatness constancy over full Arc motion of the gantry | ± 1%
5.2.6 | Beam Symmetry over full Arc motion of the gantry | ≤ 2% difference with respect to static beam symmetry
5.2.7 | Beam Symmetry constancy over full Arc motion of the gantry | ± 1%
5.2.8 | Constancy of x-ray beam quality with full Arc motion of the gantry | ±1%
@ Test pertaining to VMAT parameters are not comprehensive therefore additional tests needs to be carried out and submitted as per manufacturers’ specifications/customer acceptance tests.
6 | Performance test results of the Unflat (Flattening Filter Free) Photon Beam [Note: (1) Following additional tests need to be carried out for operation of LINAC in Unflat (FFF) photon beam mode; (2) The tolerance value indicated is the deviation of the parameter from its baseline value except for those indicated by *. Please refer AERB protocol “Acceptance Test/QA/Un-flat beam” for Quality Assurance/Acceptance Test for Unflat (FFF) beam for further information] *The tolerance value is based on the survey conducted prior to implementation of the protocol
Sr. No. | Parameters | Energies (MV) | Collimator opening (cm2) | Base Line value | Observed value | Deviation between baseline value and observed value | Tolerance Limit | Remarks, if any
6.1 | Beam quality (TPR20/10) | 6XFFF/7UF | 10x10 |  |  |  | 2%
|  | 10XFFF | 10x10 |  |  |  | 2%
6.2 | Difference in % Surface Dose (PSD) between flat and unflat beams of similar nominal photon beam energy | 6XFFF/7UF | 10x10 |  |  |  | ≤10%*
|  |  | 20x20 |  |  |  | ≤10%*
|  | 10XFFF | 10x10 |  |  |  | ≤10%*
|  |  | 20x20 |  |  |  | ≤10%*
6.3 | Off Axis Ratio (OAR) : Cross-plane (at ±3 cm from central axis at 10 cm depth) | 6XFFF/7UF | 10x10 |  |  |  | 2%
|  | 10XFFF | 10x10 |  |  |  | 2%
6.4 | Off Axis Ratio (OAR) : In-plane (at ±3 cm from central axis at 10 cm depth) | 6XFFF/7UF | 10x10 |  |  |  | 2%
|  | 10XFFF | 10x10 |  |  |  | 2%
6.5 | Depth of dose maximum (dm) (mm) | 6XFFF/7UF | 5x5 |  |  |  | ±2 mm
|  |  | 10x10 |  |  |  | ±2 mm
|  |  | 20x20 |  |  |  | ±2 mm
|  | 10XFFF | 5x5 |  |  |  | ±2 mm
|  |  | 10x10 |  |  |  | ±2 mm
|  |  | 20x20 |  |  |  | ±2 mm
6.6 | Constancy of Percentage Depth Dose (PDD) at 10 cm depth | 6XFFF/7UF | 5x5 |  |  |  | 2%
|  |  | 10x10 |  |  |  | 2%
|  |  | 20x20 |  |  |  | 2%
|  | 10XFFF | 5x5 |  |  |  | 2%
|  |  | 10x10 |  |  |  | 2%
|  |  | 20x20 |  |  |  | 2%
6.7 | Constancy of symmetry (%) | 6XFFF/7UF | 20x20 |  |  |  | 2%
|  | 10XFFF | 20x20 |  |  |  | 2%
6.8 | Separation between IPL & IPR (cm) [Refer Fig.E.1 in sheet 2] | 6XFFF/7UF | 20x20 |  |  |  | ±2 mm
|  | 10XFFF | 20x20 |  |  |  | ±2 mm
6.9 | Lateral width at 90% dose level (X90%) (cm) [Refer Fig.E.2 in sheet 2] | 6XFFF/7UF | 20x20 |  |  |  | ±2 mm
|  | 10XFFF | 20x20 |  |  |  | ±2 mm
6.1 | Lateral width at 75% dose level (X75%) (cm) [Refer Fig.E.2 in sheet 2] | 6XFFF/7UF | 20x20 |  |  |  | ±2 mm
|  | 10XFFF | 20x20 |  |  |  | ±2 mm
6.11 | Lateral width at 60% dose level (X60%) (cm) [Refer Fig.E.2 in sheet 2] | 6XFFF/7UF | 20x20 |  |  |  | ±2 mm
|  | 10XFFF | 20x20 |  |  |  | ±2 mm
6.12 | Radiation beam Penumbra (Pb-Pa) (mm) [Refer Fig.E.1 in sheet 2] | 6XFFF/7UF | 20x20 |  |  |  | ±2 mm
|  | 10XFFF | 20x20 |  |  |  | ±2 mm
7 | Performance test results of the MV imaging system (Note: Following tests need to be carried out for imaging system capable LINAC in addition to the tests for other treatment modality) *Applicable for cone beam CT
Sr. No. | Parameters | Tolerance Limit | Observation ( Write “Yes/No or give values of the measurement”, as the case may be) | Observation is meeting the requirments | Remarks, if any
7.1 | Collision interlocks provided and working
7.2 | Positional/re-positional accuracy of imager | ≤ 1 mm
7.3 | Coincidence of centre of imaging system and isocentre (vertical gantry) | ≤ 1 mm
7.4 | Geometric distortion* | ≤ 1 mm
7.5 | Spatial (high contrast) resolution | Manufacturer’s specified tolerance |  |  | Manufacturer specified value…….
7.6 | Low Contrast resolution | Manufacturer’s specified tolerance |  |  | Manufacturer specified value…….
7.7 | Uniformity and noise | Manufacturer’s specified tolerance |  |  | Manufacturer specified value…….
7.8 | Accuracy of HU values* | Manufacturer’s specified tolerance |  |  | Manufacturer specified value…….
7.9 | Constancy of HU values* | Manufacturer’s specified tolerance |  |  | Manufacturer specified value…….
8 | Safety aspects of Accelerator installation
Sr. No. | Parameters | Observation ( Write “Yes/No or give values of the measurement”, as the case may be)
8.1 | Radiation Warning legend and symbol at the treatment room entrance and at the entrance of the controlled and supervised areas is provided
8.2 | Radiation warning light at/or near the entrance to the treatment room is provided and working
8.3 | Audiovisual monitor provided and working

## Sheet2
Figure E.1: Schematic diagram for determining inflection point and penumbra | Figure E.2: Schematic diagram for determining degree of unflattenedness