## Sheet1
DO NOT MODIFY THE WORK SHEET. FILL UP THE BLUE CELLS ONLY
Name of Medical Physicist responsible for the measurements
Date of completion of the measurements
1 | General information of the Simulator Units
1.1 | Type of the unit (SIM / SIM-CT)
1.2 | Manufacturer
1.3 | Model
1.4 | Sr. No.
1.5 | Imaging capabilities (Radiography, fluoroscopy, Sim-CT)
1.6 | Imaging device (Image intensifier tube (IIT) / flat panel detector)
1.7 | Reticule provided (Yes/No)
1.8 | Range of focus to axis distance (FAD) in cm
1.9 | Maximum Tube Voltage in (kV)
1.1 | Maximum Tube Current in (mA)
1.11 | Type of Anode (Rotating / Stationary)
1.12 | Focal spot size
1.12.1 | Small |  | ___ mm x ___ mm
1.12.2 | Large |  | ___ mm x ___ mm
1.13 | Type of Rectification (Full Wave / Half Wave/ Multi pulse/Other (specify))
1.14 | Filtration
1.14.1 | Inherent (mm of Al)
1.4.2 | Added (mm of Al)
1.14.3 | Total (mm of Al)
1.15 | Range of movement of the detector panel in mm
1.15.1 | Lateral
1.15.2 | Longitudianal
1.15.3 | Vertical
1.16 | Indication of Delineated Radiation field size:
1.16.1 | at 80 cm FAD |  | Max:­­­_______ cm x cm & Min:_______cm x cm
1.16.2 | at 100 cm FAD |  | Max:­­­_______ cm x cm & Min:_______cm x cm
2 | Performance test results of the Simulator unit
Sr. No. | Parameters | Tolerance Value, if any | Observation (Write Yes/No or give values of measurment as the case may be) | Observation is meeting the requirments | Remarks, if any
2.1 | Hand held control (pendant) provided and working satisfactorily
2.2 | Displays and Safety Interlocks (e.g. collision interlock, emergency interlock) provided and working satisfactorily
2.3 | Accuracy of reticule scale | 1 mm
2.4 | Field light intensity satisfactory
2.5 | Optical Field Size (delineator) Accuracy at different FAD | 2 mm
2.6 | Beam alignment | ≤ 1.5°
2.7 | Coincidence between optical field and radiation field (both small & large focal spot ) | 1 mm for 3x3 to 20x20 cm2 & 0.5% for >20x20cm2
2.8 | Optical beam and collimator axis coincidence | 1 mm dia
2.9 | Shift in center of optical field due to minimum to maximum vertical motion of couch | 2 mm
2.1 | Shift in center of optical field due to rotational motion of the couch from -900 to 900 position | 2 mm
2.11 | Shift in center of the radiation field at the isocenter when changing from one focal spot to another. | 0.5 mm
2.12 | Shift in mechanical Isocentre due to Collimator rotation | 2 mm dia
2.13 | Shift in mechanical Isocentre due to Gantry rotation | 2 mm dia
2.14 | Shift in mechanical Isocentre due to couch rotation | 2 mm dia
2.15 | Coincidence of mechanical and radiation isocenter | 2 mm dia
2.16 | Coincidence among the Mechanical front pointer, optical distance indicator (ODI) and lasers | 2 mm
2.17 | Accuracy of ODI (60 - 120 cm) | 2 mm
2.18 | Parallelism of delineators | 1°
2.19 | Orthogonality of delineators | 90±1°
2.2 | Symmetry of delineators | ±1mm
2.21 | Field overlap for Gantry angles of 00 & 1800 and 900 & 2700 | ±1mm
2.22 | Reproducibility of delineated radiation field | 1 mm
2.23 | Angular rotation limits of the table top | ±180°
2.24 | Couch isocentric rotation limits | ±90°
2.25 | Minimum Level of table top above floor | ≤80 cm
2.26 | Longitudinal motion of couch | ≤70 cm
2.27 | Vertical motion of couch from isocentre | ˃40 cm below isocentre and ˃3 cm above isocentre
2.28 | Lateral motion of couch | ± 20 cm
2.29 | Accuracy of vertical motion of couch | 2 mm
2.3 | Accuracy of lateral motion of couch | 2 mm
2.31 | Accuracy of longitudinal motion of couch | 2 mm
2.32 | Accuracy of gantry angle | ±0.50
2.33 | Accuracy of collimator angle | ±0.50
2.34 | Accuracy of couch angle | ±0.50
2.35 | Table top sag at isocenter when loaded with 135kg distributed over 2m through isocentre | 2 mm
2.36 | Maximum horizontal shift in the projection of the cross wire due to vertical movement of the couch through 20 cm when loaded with 135kg distributed over 2m through isocentre | 2 mm
2.37 | Maximum shift in the projection of the cross wire due to isocentric rotation of the couch when loaded with 135kg distributed over 2m through isocentre | 1mm
2.38 | Accuracy of kVp | ± 5 kVp (≤10%) whichever is lesser
2.39 | Accuracy of tube current | ±20%
2.4 | Accuracy of timer/irradiation time | ±10 %
2.41 | Accuracy of current time product (mA.s) | ±10 %
2.42 | Linearity of mA station (Co-efficient of Linearity) | ≤ 0.1
2.43 | Linearity of timer (Co-efficient of Linearity) | ≤0.1
2.44 | Output reproducibility (Coeff. of Variation) | ≤ 0.05
2.45 | Focal spot size (f)
2.45.1 | for f< 0.8mm | 0.5f
2.45.2 | for 0.8 mm £ f £ 1.5 mm | 0.4f
2.45.3 | for f > 1.5 mm | 0.3f
2.46 | Total filtration (mm of Al equivalent) of the x-ray tube
2.46.1 | for kV ≤ 70 kV, | 1.5 mm Al
2.46.2 | for 70 < kV≤ 100 | 2 mm Al
2.46.3 | for kV> 100 | 2.5 mm Al
2.47 | Radiation leakage level at 1m from the tube housing | <1 mGy in 1 hr
2.48 | Low contrast resolution | 3.0mm at 10% contrast difference
2.49 | Spatial resolution | Within 10% of specified value
3 Additional performance test to be carried out for Simulator having SIM-CT option
Sr. no. | Parameters | Tolerance | Observation | Observation is meeting the requirments | Remarks, if any
|  | Limit | ( Write “Yes/No/ or give values of the measurement”, as the case may be)
3.1 | Slice thickness accuracy (sensitivity profile width) | ±1mm for slice thickness > 2 mm
3.2 | CT dose index (CTDIW ) | ±20% of quoted value
3.3 | Laser accuracy | ±1mm
3.4 | CT number uniformity | ±4 HU for center and ±2 for other location
3.5 | CT number Accuracy | ±4 HU
4 | Safety aspects of Simulator installation
4.1 | Radiation Warning legend and symbol at the entrance provided: (Yes/No)
4.2 | Radiation warning light at the entrance provided and working: (Yes/No)
4.3 | Audiovisual monitor provided and working: (yes/No)
4.4 | Whether entrance doors lined with 2mm lead are provided: (yes/No)
4.5 | Whether viewing window with 2 mm lead equivalence is provided: (yes/No)